## Slide 1
Progress of heavy-ion accelerator(RAON) including Linac commissioning
Yeonsei Chung on behalf of Rare Isotope Science Project
ASSCA2022, Feb. 14, 2023

## Slide 2
001. Project Overview & Status
002. Accelerator Systems
004. Beam Instrumentation & commissioning
005. Summary and Outlook
003. RI & Experimental Systems

## Slide 3
Part 1.
Project Overview

## Slide 4: Rare Isotope Science Project (RISP)
Providing high intensity RI beams by ISOL and IF
ISOL: direct fission of 238U by 70 MeV proton
IF: 200 MeV/u 238U (intensity: 8.3 pμA)
Providing high quality neutron-rich beams
e.g., 132Sn with up to 250 MeV/u,
up to 109 particles per second
Providing More exotic RI beam
production by combination of ISOL and IF
Budget: KRW 1,518 billion (US$ 1.32 billion, 1$=1,146krw)
- accelerators and experimental apparatus : 522.8 billion won
- civil engineering & conventional facilities : 996 billion won (incl. site 357 billion won)
Period: 2011.12 ~ 2022.12(1st Phase)
2023.01 ~ (2nd Phase)
Goal: To build a heavy ion accelerator complex RAON,
for rare isotope science research in Korea.
* RAON - Rare isotope Accelerator complex for ON-line experiments
Construction of research and support facility to ensure the stable operation of the heavy-ion accelerator, experiment systems, and to establish a comfortable research environment
※ Accelerator and experiment buildings, support facility, administrative buildings, and guest house, etc.
Development, installation, and commissioning of the accelerator systems that provides high-energy (200MeV/u) and high-power (400kW) heavy-ion beam
System Installation Project
Facility Construction Project

## Slide 5
Campus Layout
Assembly Bd.
SRF Test Bd.
High Energy B
IF Separator / High Energy A
HQ Office
Guest House
Info. Center
Cryogenics Bd.
ISOL
SCL
Utility Bd.
Electricity Bd.
Control Center
Low Energy A,B
Waste Strorage
Reservation
Campus Area : 952,066㎡ (including the reservation area of 144,640㎡)
Period : 2014 ~ 2021(8 years)
Cost : 996 billion won (incl. 357 billion won for land)
Building Area : 76,259㎡(11 Bds)
Total Bd. Area : 116,252㎡
Constructor : POSCO Consortium(11 companies)

## Slide 6: RAON Layout
● SCL1 has been decided to postpone
: SCL3 is going to be taking a role of SCL1 in the early operation
IF System
BIS
μSR
HPMMS/
CLS
SCL2
ISOL System
LAMPS
Cyclotron
SCL3
SCL1
ECR IS
KOBRA
NDPS
Control
Center
Utility
HRS

## Slide 7
RAON will provide access to unexplored regions of the nuclear chart
KOBRA
IF
ISOL
ISOL+IF
RIBs at RAON

## Slide 8
View of Construction Site
<`20.10>
<`17.01>
<`18.01>
<`19.01>
<Assembly Bd.>
<SRF Test Bd.>
<High Energy Exps.>
<ISOL>
<SCL3>
<Low Energy Exps.>
<Control Center>
● View of Construction Place (`20.10)

## Slide 9
Assembly Bd.
Control Center
SRF Test Bd.
Utility Bd.
Electricity Bd.
HQ Office Bd.
Conventional Facilities

## Slide 10
SCL3
Low Energy A/B
SCL2
IF/ High Energy A
High Energy B
ISOL
Accelerator Building, Experimental Halls
SCL3
Bending Section
SCL3-gallery
10

## Slide 11
Part 2.
Accelerator Systems

## Slide 12
Accelerator Systems

## Slide 13
MEBT
RFQ
LEBT
SC ECRIS
ECRIS
Injector System
Two ECR-IS on high voltage platforms
14.5 GHz ECR ion source
28 GHz superconducting ECR ion source
LEBT (E = 10 keV/u)
10 keV/u, Dual bending magnet
Chopper & Electrostatic quads, Instrumentation
RFQ (E = 500 keV/u)
81.25 MHz, Transmission Eff. ~98%
CW RF Power 94 kW (SSPA: 150 kW)
MEBT (E = 500 keV/u)
Four RF bunchers (SSPA: 20, 15, 2×4 kW)
Simple quadrupole magnets, Instrumentation
Installation completed and beam commissioning from October, 2020

## Slide 14
Control System
Control Center
Main Control Room
Data Storage System
Local Control System
Beam Diagnostics Control System
Integrated Control System
Fast Protection System
EPICS IOC
Controllers
SCL3 Control System(43)
Timing System

## Slide 15
Cavity String
With cryogenic piping
Vacuum leak rate
Particle count inside cavity “0”
Top loading to cryostat
HWR B type
HWR Cryomodule Assembly

## Slide 16
Clean Assembly @ Accelerator Tunnel
(Cryomodule + Warm section) + (Cryomodule + Warm section)
Cryomodule & Warm section is clean assembled in the clean booth@tunnel
Total Particle counts(size=0.5um above/10 mins) were less than 30 counts

## Slide 17
Superconducting Linac, SCL3 Tunnel and Gallery
QWR & HWR Cryomodule
Cryogenic Distribution to Cryomodule
CM/Cryogenic Control Rack and SSPA
Clean beam line assembly
Installation completion and ready for beam commissioning in 2021

## Slide 18
Accelerator Systems – SCL2
SCL21(SSR1)
∙ 1st prototype was performed by TRIUMF(2019.6)
Balloon type(less multipacting, better mechanical characteristics)
Accelerating gradient over 8.7MV/m, Multipacting below 2MV/m
∙ 2nd prototype with domestic vendor
2nd prototype by domestic vendor
SCL22(SSR2)
∙ 1st prototype of cavity
Designed by RISP, fabricated by domestic vendor
Balloon type, Deep drawing (depth ~280mm)
∙ 2nd prototype with IHEP
Cylindrical type
RISP design(balloon type)
1st prototype by TRIUMF
∙ SSPA (8, 20 kW): 2020.01~2022.12
Requirement for RF power rating
∙ HPRF transmission line
3-1/8”, 4-1/16” rigid coaxial
∙ LLRF control system (2020.04~2023.12)
∙ RF reference line (81.25MHz) : 2022.12
SSR1
SSPA Prototype
RF System

## Slide 19
3 pits for VT; up to 3 cavities per pit
3 HT bunkers
To cover all RAON cavities - QWR (82.125 MHz), HWR (162.5 MHz) and SSR1 & 2 (325 MHz)
On-site SRF
test facility
SRF Test facility

## Slide 20
Frequency (MHz) | Optimum 𝛽 | Eacc (MV/m) | Vacc (MV) | Q0 | # of cavity | # of module | Op. temp (K)
81.25 | 0.047 | 6.1 | 1.06 | 1.3e+8 | 22 | 22 | 4.5
Target
QWR Cavity VT test results

## Slide 21
Target
1.5e+9 @6.6 MV/m
The cavities below the target include bare cavities as R&D phase and all were re-processed after 1st VT
Frequency (MHz) | Optimum 𝛽 | Eacc (MV/m) | Vacc (MV) | Q0 | # of cavity | # of module | Op. temp (K)
162.5 | 0.12 | 6.6 | 1.46 | 1.5e+9 | 106 | A/B = 14/19 | 2.05
# of fabrication (@10/12/2020) |  |  |  |  | 53 | A : 8, B : 1 | 
HWR Cavity VT test results

## Slide 22
81.15 MHz (SEL)
81.25 MHz (SEL)
162.4 MHz (SEL)
Test and confirmation with SEL(Self-Excited Loop) mode(±100 kHz) & GDL(Generator Driven Resonator) mode; GDR mode@nominal operation
Realized, LLRF and control system via EPICS
LLRF Latency Test
QWR Cavity Control Test@SRF
CSS OPI
Eacc=6.11 MV/m (±0.49%), Phase=50o (±0.14o)
LLRF System
162.5 MHz (SEL)

## Slide 23
RF Control Test(Amplitude/Phase feedback with Tuner control)
RF Amplitude, Phase Control
Target RF amplitude, 6.6 MV/m(±0.13%) < 1%, Phase=100o (±0.63o) < 1 deg
Tuner operation threshold : +/- 5deg
Control bandwidth: 90~160 Hz
RF Control Test of HWR#26

## Slide 24
Cryo-plant

## Slide 25
Cryogenic Distribution Systems
SCL2
SCL3
HWRB(19)
HWRA
(13 +2)
QWR(11)
SSR1(23)
SSR2(23) + Extension (2)
LTS(13)
SCL2
SCL3
Layout of cryogenic distribution system @ SCL3 and SCL2

## Slide 26
4.5 K reached
Cavities : Cool-down
Cavities : QWR Cool-down
Cavities : HWR Cool-down
End
Oct. 26th
QWR CM1, 2
HWR B CM19
Cooling down cryogenic distribution system, thermal shields of all CM with SCL3 cryo-plant, simultaneously.
The First Cool-Down Curve of SCL3
From Sep. 7th

## Slide 27
Part 3.
RI & Experimental Systems

## Slide 28
RI & Experimental System
Low Energy Exp. Bldg
Ultra-low Exp. Bldg
High Energy Exp. Bldg
CLS
HPMMS
Neutron Facility
KOBRA
mSR
LAMPS
ISOL
IF
Bio-medical facility

## Slide 29
ISOL System
TIS, EBIS, RFQ-CB, beam line were tested with SI(Cs) beam, completed in 2021
MMS/MR-TOF
EBIS Charge Breeder
ISOL Beam Line
ISOL Target Room

## Slide 30
ISOL System
• Driver beam : proton 35<K<70 MeV, up to 70 kW
• Target : UCx, MgO, BN, CaO, BeO, SiC, etc
• Ion Source : Surface, RILIS, Plasma
• RIB : 6< A < 250, 10<K< 80 keV, 10^8 pps(Sn), >90% purity @Exp. • incident to RFQ of Post accelerator 10 keV/u
• full remote maintenance system with TIS modularization
- ISOL beam lines including sub-systems
are commissioned with Cs beam in 2021
- RI beam commissioning using SiC target (Q1 2023)

## Slide 31
Magnetic current (A)
ISOL Beam Test in 2021

## Slide 32
KoBRA Spectrometer commissioning in 2021
Alpha source (241Am)
2021.10.15
2021.10.18
F1 PPAC
2021.10.19

## Slide 33
Other Experimental Systems
All exp. systems are installed and machine-commissioned by 2022
(Cyclotron)
(muSR)
(MMS)
(LAMPS)
(BIS)
(IF Separator)
(NDPS)

## Slide 34
Beam Instrumentation & Commissioning
Part 4.

## Slide 35
Beam Diagnostics Functions
Initial commissioning & Component tuning
FC, WS, BV, ACCT, BPMs(position, phase, TOF)
Beam specification : Ar(9+), ~30eμA, 100μs, 1Hz
During operation(On-line)
Monitor beam transport and acceleration function
BPM : Beam position and phase
BCM : Beam current and transmission(RFQ, SCL3, P2DT, SCL2)
BLM : Beam Loss and link to Machine Protection
Beam current is 1mA ~ 1μA
Commissioning and during operation(On-demand)
1-D, 2-D profile distribution (WS, EM)
Bunch length (Stripline FFC)
Beam Attenuator @LEBT

## Slide 36
Injector Instrumentation
LEBT
Wire scanner - 5 sets (4 installed)
Beam Viewer – 3 sets (2 installed)
Bergoz ACCT – 1 set
Faraday Cup – 5 sets (4 installed)
beam Attenuator – 2 sets (1/100)
2-D Allison Scanner – 2sets (1 installed)
MEBT (Installed)
Wire scanner - 4 sets
Beam Viewer – 1 set
Bergoz ACCT – 2 sets
Faraday Cup – 2 sets
Stripline FFC – 1 set
BPM – 6 sets
4.

## Slide 37
Layout of RAON Beam Diagnostics

## Slide 38
Beam Position Monitor(BPM)
Measurements : Beam Position, Phase, relative Intensity
SCL3(Installed) : Button-type BPM, Curved type 20mmx30mm
P2DT(Installed) : Large button-type BPM(Bending section), Button-type BPM
SCL2(Fabricated) : Stripline BPM, Curved type 25mmx150mm

## Slide 39
Beam Diagnostics at SCL3-Warm Section
BPM
BPM and BDC inside of Magnetic Quadrupole Doublet
Installation (& alignment) procedure of BDC/BPM/Beam Pipe assembly (particle-free with pure N2 purging) in between two cryomodules is well defined & followed
Ring Type(ID=36 mm, Nb) BLC inside BDC
- Directly collect halo/stray beam
Cryomodule
Beam
Diagnostic
Chamber
Quadrupole
Magnet
BPM

## Slide 40
Beam Loss Monitor(BLM)
Superconducting Linac : 1 BLM(BLC, PD)/Warm section
Bending Section(P2DT) : Plastic Scintillator, Ion Chamber
DBCM(ACCT networks) for Primary MPS
PC or IC
ACCT
ACCT
ACCT
ACCT
ACCT
ACCT
ACCT
ACCT
ACCT
ACCT
 | PC(IC) | PD | BLC | DBCM(ACCT)
QWR | - | - | ongoing | ongoing
HWR1 | - |  | ongoing | ongoing
HWR2 | - |  | ongoing | ongoing
P2DT | X | X | X | X
SSR1 |  | X | X | X
SSR2 |  | X | X | X

## Slide 41
Diagnostics DAQ System
AMC-pico8
Control board
BPM electronics
cRIO for ACCT
BPM
153 modules fabricated; SCL3(60), P2DT(28), SCL2(65)
supports timing, trigger, interlock, postmortem, EPICS of RAON
Current (WS, FC, BLM, BLC) - μTCA system
CAENels AMC-pico: 42 boards, 8 ch/20 bit/pA range
Modified firmware for RAON trigger, interlock and EPICS
ACCT – Standalone 1u chassis
100 MS/sec, 14-bit 4 ADC channels
Upgrade from cRIO System(1MS/sec)
supports timing, trigger, interlock, postmortem, EPICS of RAON
Electronics for SCL3 are fully tested/calibrated

## Slide 42
Beam Physics
Start-to-end simulation: uranium (238U33+,34+→77+,78+,79+,80+,81+)
SCL3
SCL2
Physics application tools
BIPAM (Beam Input Parameter And Matching)
CAPS (Cavity Amplitude and Phase Scan)
SCL2 extension
On-line orbit correction
Transverse beam size
Longitudinal beam size

## Slide 43
Beam parameter measurements (Allison scanners, wirescanners)
measuring initial beam parameters (fitting beam sizes of wirescanners)
controlling optics when needed
do matching to RFQ
Emittance measurement (Allison scanner, quad scan)
Beams: Ar9+ (~30A), Ar8+ (~47A)
LEBT beam parameters
Beam emittance
(Allison scanner)
 | X | Y
Allison | 0.048 | 0.067
Quad Scan | 0.041 | 0.038
emittance comparison
Beam emittance
(quad scan)
Beam commissioning (LEBT)

## Slide 44
Beam commissioning (LEBT)
[ Ar8+ ]
[ Ar9+ ]
- modelling of initial beam using wire scanner and Allison data
* Comparing WS data with TRACK simulation
* First wire scanner is located before Allison scanner
TRACK (blue)
Measured(red)

## Slide 45
- orbit correction
- beam tracking (TRACK, DYNAC codes)
LEBT orbit correction
TRACK (beam dynamics code)
LEBT Orbit Correction

## Slide 46
RFQ RF set-point (Ar9+, Ar8+):
beam transmission measured using MEBT ACCT2
Fitting against model
* Measured transmission = 94 % (simulation = 98%)
Cavity RF power: 51.5 kW (Design ~39.1 kW (20% margin))
RFQ RF set-point
(error bar: 3)
RFQ
MEBT
measurement
model
RFQ Beam Transmission

## Slide 47
MEBT beam parameter measurements
using beam sizes from wirescanners
measure initial beam emittances & parameters
can do matching to SCL3
MEBT buncher RF set-point
using phase scan technique
determines RF amplitude & phase
RF set-points of 4 bunchers obtained
Beam Energy is 507 keV/u
(TOF, 1st and 2nd BPMs in MEBT)
ex=0.052
ey=0.059
Medium Energy Beam Transport (MEBT)

## Slide 48
- 10% beam duty operation: 96 minutes, 10Hz, 10msec (2021.12.07.)
* Injector transmission > 94%
- MEBT beam emittance measurement based on quad scan
x = 0.064 mm-mrad
y = 0.049 mm-mrad
Injector beam transmission
MEBT quad scan
emittance measurement
Injector Transmission
TRACK (blue)
Measured(red)

## Slide 49
PEEK
Ta plate
Amplifier ( 43 dB Gain )
Bandwidth 300 kHz ~ 14 GHz
Frequency component < 3 GHz
With 0.13 ns bunch length
Oscilloscope (4 GHz, 25 GSPS)
Fabrication of Stripline type Fast Faraday Cup
Semi-rigid SMA cable in vacuum
PEEK insulator
Ta plate in front of FFC
Bolting at irregular position
Ar 8+, 50 uA, at the end of MEBT (4 bunchers)
macro pulse commissioning beam
Expected peak amplitude was ~ 4 mV
RF amplifier and oscilloscope prepared, considering frequency component
Bunch Length Measurement, Fast-FC

## Slide 50
debunching
bunching
Phase scan of Buncher 4
Measured bunch Length at MEBT
Measured single bunch length was with Gaussian fit.
Shows good signal repeatability for measuring bunch trains.
Overlapped signal of 325 bunches in 4 us is shown.
Bunching and debunching was observed by rebuncher in MEBT.
Bunch Length Measurement

## Slide 51
Measurements to perform
QWR/HWR RF set-point (phase scan, BPM)
Beam energy (phase scan, TOF with BPMs)
Beam current, transmission (ACCT, FC)
Orbit correction (BPM & dipole steerer)
Beam transverse profile (wirescanner)
Transverse matching (wirescanner)
Beam emittance (X and Y) : beam size fitting, quad scan
Physics applications for commissioning
BIPAM (Beam Input Parameters And Matching)
CAPS (Cavity Amplitude and Phase Scan)
Orbit correction application
Emittance data analysis & wirescanner data analysis
Linac(SCL3) Commissioning

## Slide 52
  Ar9+ beams accelerated by QWR #1~#5 on the 7th of Oct, 2022
Beam Energy 0.701 MeV/u
The 1st SCL3 Beam Commissioning (Oct. 7, 2022)

## Slide 53
Phase scan with BPMs (Time of Flight measurement)
CAPS (Cavity Amplitude and Phase Scan)

## Slide 54
The 2nd SCL3 Beam Commissioning (Dec. 16, 2022)
  Ar9+ beams accelerated by QWR #1~#22 on the 16th of Dec, 2022
Beam Energy 2.472 MeV/u
Final Results :
2.457 MeV/u (Measured)
2.444 MeV/u (TRACK)
QWR#3, #20 excluded

## Slide 55
SCL3 Beam Commissioning Plan
Finished 2 K cool-down of SCL3 on Jan. 11, 2023.
Commissioning of SCL3 to be finished until March 2023
From March 2023, beams on KoBRA spectrometer

## Slide 56
Summary & Outlook
Part 5.

## Slide 57
Injector beam commissioning was carried out, achieving machine setting and key measurements :
measured beam parameters (energy, emittance, Twiss parameters, beam sizes etc)
capable of controlling LEBT and MEBT beam optics freely as needed
achieved beam transmission of 95% max (routinely > 90%)
machine verification including diagnostics devices
RAON is preparing for Linac(SCL3) beam commissioning
1st/2nd beam commissioning using 22 QWR CMs in SCL3 were successfully done.
ECR → SCL3 → KoBRA using Ar-40(9+) in Q1/Q2 of 2023
Cyclotron → ISOL for RI beam extraction is expected in Q1 of 2023
Plan for SIB/RIB experiments
RIBs from ISOL to be injected into SCL3 in Q4 of 2023
SIB experiments(ECR→SCL3 → KoBRA/NDPS) will be carried out in Q4 of 2023
RIB experiments in ISOL beamline using Al isotope beams
Summary & Outlook

## Slide 58
감사합니다.
Thank you